DETERMINATION OF GLYCOLS IN BIOLOGICAL SPECIMENS BY GAS CHROMATOGRAPHY-MASS SPECTROMETRY

J.-P. Goullé, V. Gembus, Ch. Lacroix

# Groupe Hospitalier, BP 24, 76083 Le Havre, France

A simple extraction and derivatization procedure for the analysis of eight glycols (ethylene gycol = EG; diethylene glycol = DEG; 1,3-propanediol = 1,3-PD; 1,2-propanediol = 1,2-PD; 1,2-butanediol = 1,2-BD; 2,3-butanediol = 2,3-BD; triethylene glycol = TEG; hexylene glycol = HXG) using a 20 µl ten fold diluted serum or blood sample is described. Following deproteinisation with acetonitrile, derivatization to its mono, di or tri TMS derivative, the glycols were detected using gas chromatography-electron ionization mass spectrometry equipped with a split-splitless inlet and a DB 5MS column in the scan mode from 40 to 500 u. Gamma-hydroxybuturate d6 (GHB-d6) was used as the internal standard. With one quantitation ion and two confirmation ions, the limit of detection and the limit of quantitation in serum ranged from 0.7 mg/l for EG to 8,5 mg/l for TEG and from 1.3 mg/l for EG to 18.2 mg/l for 1,2-PD, respectively. A linear response was observed over the concentration range from 1-18 to 800 mg/l for serum and blood. Coefficients of variation for both intra-assay precision and inter-assay reproductibility ranged between 1.9 % for TEG to 4.9 % for 1,2-PD (11.8 % for HXG) and 3.5 % for DEG to 9.0 % for 2,3-BD (20.4 % for HXG) at the 400 mg/l serum level, respectively. The method was applied to plasma and whole blood.

Keywords: glycols, gas chromatography mass spectrometry

(jgoulle@ch-havre.fr)