IMMUNOASSAY SCREENING OF LYSERGIC ACID DIETHYLAMIDE (LSD) IN URINE AND ITS CONFIRMATION BY HPLC WITH FLUORESCENCE DETECTION AFTER IMMUNAFFINITY EXTRACTION

T. Grobosch, U. Lemm-Ahlers

Berliner Betrieb für Zentrale Gesundheitliche Aufgaben (BBGes) FB Klinische Toxikologie und Pharmakologie

Nearly 4000 urine specimens, mainly from psychiatric patients or drug abusers were screened for LSD using the CEDIA DAU LSD assay. LSD positive samples (n = 48, cut-off: 500 pg/mL) were confirmed by HPLC with fluorescence detection (HPLC-FLD, column: LiChrospher 100 RP 8 cartridge, 150 x 4.0 mm, 5 µm; mobile phase: acetonitrile - 50 mM potassium phosphate buffer pH 2.3 (34 : 66, v/v), ex/em: 315/420 nm, linearity: 50-500 pg/mL, LOQ: 50 pg/mL) after immunaffinity extraction of 3 mL of urine (IAE, Microgenics LSD ImmunElute resin and PolyPrep disposable extraction column). The IAE/HPLC-FLD combination was rapid, easy to perform and reliable.

The CEDIA DAU LSD positive specimens were additional tested by DPC Coat-a-Count RIA in order to compare both immunoassay screening methods with HPLC confirmation. Complete agreement in the qualitative mode was observed between the DPC RIA assay and HPLC-FLD results when a cut-off concentration of 500 pg/mL was used. In general, both immunoassays tend to give higher readings for LSD than HPLC-FLD.

Samples which were CEDIA DAU LSD positive but not confirmed by HPLC-FLD were further investigated for interfering compounds using REMEDI HS drug profiling system. REMEDI HS analysis identified eighteen compounds (parent drugs and metabolites) which may probably cross-react with the CEDIA DAU LSD assay: ambroxol, prilocaine, melperone, pipamperone, diphenhydramine, metoclopramide, methaqualone, amitryptiline, doxepine, atracurium, bupivacaine, doxylamine, lidocaine, mepivacaine, promethazine, ranitidine, tramadol. As the cross-reaction of all these parent drugs was < 0.001 %, only their metabolites could have cross-react. Further studies are needed for confirmation of this hypothesis.

Referee's comment:

After revision (the points were discussed by phone with the authors and already performed) the abstract can now be accepted as oral presentation in the CT session.

(grobosch.tox@bbges.de)