ROADSIDE DRUG TESTING: THE RESULTS OF THE ROSITA PROJECT.

A.Verstraete for the Rosita consortium.

Laboratory of Clinical Biology– Toxicology, Ghent University Hospital, Belgium

# Aims

According to several experts, a better enforcement of the legislation on driving under the influence of drugs can be reached by training of the police officers in drug recognition and by the availability of reliable roadside drug tests. One of the objectives of the Rosita project was to evaluate the use of onsite drug tests at the roadside.

## Materials and methods

During the Rosita project, 2968 subjects were tested in 8 countries. In 6 countries (Belgium, Finland, Germany, Italy, Norway and Spain), the tests were evaluated at the roadside or in a police station. Onsite immunoassays were used for the detection of drugs in urine, oral fluid and/or sweat. Confirmation analyses on blood, urine, oral fluid and/or sweat, mainly by GC/MS, but in some cases by GC/ECD and HPLC-DAD, were performed in the participating laboratories.

### Results

Police officers liked having the tools to detect drugged drivers, and they were very creative in finding solutions to the practical and operational problems they encountered. Onsite drug testing gave police confidence, saved time and money.

Police officers had no major objections to collecting specimens of body fluids. In the majority of the participating countries, oral fluid was the preferred specimen. Obtaining a urine specimen was no problem if the necessary facilities (e.g. a sanitary van) were available, but was more difficult if this was not the case.

Some onsite urine tests (Rapid Drug Screen, SYVA Rapidtest, Dipro Drugscreen 5, Triage) yielded good results (accuracy >95%, sensitivity and specificity >90% compared with reference methods), but none gave a good result for all assays. For the detection of phenethylamines (including ecstasy), the combination of an amphetamine and methamphetamine test often gave excellent results.

Oral fluid and sweat are promising specimens and in some cases are better than urine but more development of the onsite tests will be needed. The sampling of an oral fluid or sweat specimen was well accepted by drivers. The onsite tests that were evaluated (Drugwipe, Cozart Rapiscan and Avitar Oralscreen) were not sufficiently reliable (accuracy between 50 and 81 % in comparison to blood). Progress is needed for sampling, duration of the test, sample volume and reliability. For cannabis and benzodiazepines, the present generation of tests is not sensitive enough. New prototypes of onsite oral fluid drug tests have been announced, and new studies will be needed for their evaluation.

### Conclusion

Roadside drug tests were considered to be very useful. In the future, the use of oral fluid seems the most promising, but the presently available tests are not satisfactory. In the meantime, urine tests can be an acceptable alternative.

Keywords: Onsite drug testing; Urine; Oral fluid; Sweat; Blood

(alain.verstraete@rug.ac.be)