# INTERFERENCE OF PHOLCODINE WITH RIA AND GC-MS OPIATE ANALYSES

C. Meadway

The Regional Laboratory for Toxicology, City Hospital NHS Trust, B18 7QH, Birmingham, United Kingdom

Pholcodine (3-morpholinoethylmorphine) is an antitussive agent, available in over the counter preparations in many countries. Structural similarities between pholcodine and morphine can lead to difficulties when interpreting analytical data produced during opiate analysis.

The manufacturers of pholcodine (Pinewood Healthcare, Co. Tipperary, Ireland) state that morphine is present as a minor impurity in pholcodine Linctus.

Analysis of a 0.1 mg/mL solution of pholcodine linctus using the DPC coat a count serum morphine radioimmunoassay (RIA) determined the morphine content to be 38 ng of morphine per 1 mL of Pholcodine linctus solution. Analysis of the same pholcodine solution using GC-MS identified the average morphine content to be 39 ng of morphine per 1 mL of Pholcodine linctus solution.

The data described above suggests that pholcodine may cross react with the DPC RIA, leading to problems if RIA is used as a confirmatory technique. Equally morphine may be produced analytically by cleavage and subsequent derivatisation of pholcodine in the injection port of GC-MS systems.

DPC have subsequently confirmed that the coat a count serum morphine RIA may cross-react with pholcodine if it is present in high concentrations.

The study was extended further to included data from volunteers following the administration of a standard dose of pholcodine (10 ml of 1mg/mL linctus). The maximum urinary pholcodine concentration measured in a spot urine by GC-MS was 3089 ng/mL at 8 hours post dose. Morphine was detectable at 2 and 4 hours post ingestion at concentrations of 151 ng/mL and 214 ng/mL respectively, but was not detected at 8 hours post dose (GC-MS limit of detection 20 ng/mL).

It is not possible to determine from these studies whether morphine is present as a metabolite of pholcodine, a ‘contaminant’, or is produced as an analytical artifact. Care must be taken when interpreting opiate results in the presence of pholcodine, especially if immunoassay results are not confirmed.

Keywords: Pholcodine, Morphine, RIA, GC-MS

(Claire.Meadway@cityhospbham.wmids.nhs.uk)