SIMULTANEOUS DETERMINATION IN URINE OF COCAINE AND ITS MAIN METABOLITES BY GAS CHROMATOGRAPHY WITH A NITROGEN - PHOSPHORUS DETECTOR

J.F. Sánchez Bruzón1, I. Farrás Fernández2, H. Cruz Rodríguez2, B. Valera Lozano2

1Antidoping Laboratory, Institute of Sport Medicine. Havana City, Cuba.

2Institute of Legal Medicine, Havana City, Cuba.

The simultaneous detection of cocaine (COC) and its‘ metabolites, benzoylecgonine (BE), ecgonine methyl ester (EME) and ecgonine (EG) in urine, is a reliable method for confirmation of COC use.  The purpose of this paper is the validation of a qualitative and quantitative method for the simultaneous analysis of COC, cocaethylene (COCET), (underivatized) BE, EME and EG in urine as the trimethylsilyl derivatives (BSTFA) with lidocaine (LIDO) as internal standard.  The method involves solid-phase extraction (SPE) utilizing Evidex cartridges along with a gas chromatographic system equipped with a DB-1 widebore  column (15 m x 0.53 mm x 3 µm). The temperature program was as follows: 60 C (held for 2 min) to 140 C at 20 C/min, 170 C at 4 C/min, 208 C (40 C/min), 214 C at 6 C/min (held for 6 min), 244 C at 30 C/min (held for 10 min).  A nitrogen-phosphorous detector (NPD) was used with N2 as carrier gas (4 mL/min).  The calibration curve was linear from 406-4060, 224-2240, 108-1088, 116-1167, 184-1840 ng/mL for EME, EG, COC, COCET and BE, respectively.  The relative standard deviation of response factors was < 5% in all cases.  Recovery was determined to be 39, 62, 92, 85 and 92% to EME, EG, COC, COCET and BE, respectively.  The method was validated for use with both spiked and cocaine-users urine specimens.  This is the first report of simultaneous SPE of COC and its‘ primary metabolites with Evidex cartridges.

Keywords: cocaine and metabolites, urine, widebore gas chromatography, Evidex - SPE

(juanfco@antidop.inder.co.cu)