DEVELOPMENT OF A SIMPLE AND UNIVERSAL EXTRACTION PROCEDURE FOR THE SAMSHA DRUGS OF ABUSE (DOA) USING THE WATERS OASIS MCX EXTRACTION CARTRIDGE AND THE ZYMARK RAPIDTRACE SPE ROBOTICS INSTRUMENT

D. Morrison

# Waters, France

Waters Oasis® MCX cartridges contain a novel, polymeric, water-wettable, sorbent, stable from pH 0 to 14. This polymer has no free silanols to complicate retention, a 1meq/g cation exchange capacity and a reverse phase with more than three times the capacity of a C18 sorbent. The mixed bed sorbent of the MCX provides for the efficient extraction of acidic, neutral and basic drugs. Using the Oasis® MCX Cartridges our objective was to develop a simple and universal procedure for extraction of the SAMSHA DOA from urine that minimized the use of organic solvents and buffers. The following procedure was developed: After loading (1 to 3 mls Urine) onto the Oasis® MCX Cartridge, the Cartridge is washed sequentially with NaOH, Water and Hexane. No pretreatment of the Cartridge is required prior to Loading. The Cartridge is eluted with an IPA/Methylene Chloride mixture, evaporated to dryness with Nitrogen, and derivatized (except for PCP) prior to GC/MS analysis. This universal 4-step extraction procedure was validated for all of the SAMSHA DOA using the Zymark RapidTrace SPE Robotics Instrument. The Controls in the Validation Protocol were set at 40% of SAMSHA Cutoff (Low), SAMSHA Cutoff (Medium) and 2X or greater of the SAMSHA Cutoff (High). (The Cutoff for Morphine & Codeine was 300 ng/ml.) The Validation Study demonstrated CVs < 10% at the Low Control and < 5% at the Medium and High Control for all of the SAMSHA DOA. This procedure, using the Oasis® MCX Cartridge, was demonstrated to provide an efficient and highly reproducible universal extraction procedure for all of the SAMSHA DOA.

Keywords: SAMSHA Drugs of Abuse, Waters Oasis® MCX, GC/MS

(Damian_morrison@waters.com)