# SUPRESSION EFFECTS IN LC-ESI/CID-MS FOR DRUG SCREENING ANALYSIS

M. Stoertzel

Introduction: Supression effects in LC-ESI-MS screening analysis for drugs were investigated, since they might be important for discrimination of certain compounds. Six different methods which are commonly used in forensic and clinical toxicology were compared.

Methods and Instrumentation: The following sample extraction methods were used to generate the blank extracts: liquid-liquid extraction, protein precipitation, solid phase extraction (SPE) with mixed mode silica, acid/neutral and alkaline fraction with polymer- based cation-exchange column, combined with protein precipitation. A Sciex API 365 with turboionspray source was used in Q1-scan mode, extracts were analyzed with polar-RP (150x2mm, 4µm, Phenomenex), gradient elution, postcoplumn continous infusion was used for detection of suppression effects.

Preliminary Data: Supression was detected in the LC-front peak (at tr=2min of solid-phase extracts and protein-precipitated samples. Less suppression was found in the second fraction of SPE extracts (basic fraction), no suppression was detected with liquid-liquid extraction in serum.

(stoertz@ukl.uni-freiburg.de)