DETERMINATION OF NALTREXONE AND ITS MAJOR METABOLITE 6--NALTREXOL IN SAMPLES FROM HEROIN ADDICTS ON NALTREXONE IMPLANTS TREATMENT

L. Olsen, A.S. Christophersen

# National Institute of Forensic Toxicology. P.O. Box 495 Sentrum 0105 Oslo, Norway

A sensitive and specific high-performance liquid chromatographic (HPLC)  method combined with electrochemical (EC) and UV detection has been modified* and validated for the analyses of naltrexone and its major metabolite 6--naltrexol in plasma.  The analytical procedure includes liquid/liquid extraction with dichloromethane/isopropanol (19:1) after alkalinazing (pH 9), back-extraction in phosphoric-acid (0,017 M) and direct injection  on a Guard column, 10 x 2 mm, which was coupled to Spherisorb 3 ODS-2, 100 x 4,6 mm, 3m, as analytical  column, 40oC, connected to both EC- and UV-detector. Acetonitrile (28%)/  phosphoric acid (40mM)/ SDS (2mM) was used  as mobile phase, flowrate 0.8 ml min Ketobemidon was used as internal standard (IS).  The recovery was found to be approximately 80% for both naltrexone, 6--naltrexol and the IS. RSD  for intra-day and inter-day variations were approximately 5% and 10%, respectively for both naltrexone and 6--naltrexol, while  LOD was 1 ng/ml for both compounds.  The specificity of the method was tested by analysing the most commonly abused drugs. The method has been used for the analyses of plasma samples from heroin addicts on naltrexone treatment used as  implants (1 g, duration 5 – 8 weeks).  Naltrexone and 6--naltrexol concentrations varied from 1-7 ng/ml and 3 – 50 ng/ml, independent of the number of days after the last implant.

* A. F. Davidson, T.A.Emm and H. J. Pieniasczek, J. Pharm. & Biom. Anal. 14 (1996) 1717

Keywords: Naltrexone, metabolite, HPLC, heroin addicts, naltrexone implant treatment

(mette.krogh@farmasi.uio.no)